|  |  |  |  | 2026 NEW Compressor Gas Cooler |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | CCA 992C |  |  |  |  |  |  |  |  |
| انتخاب کد مورد نظر: |  |  |  |  |  |  |  | Select Required Code: |  |  |  |  |
| CCA 992C; 0 1; بدون مبدل با مبدل; 0 (None) G (Duran Glass) S (SS316) P (PTFE / PVDF); جنس مبدل; 1 2 3 4; تعداد مبدل; 1=150 2=250; فلوی گاز Nl/hr; 0 (None) 1 (SR25 Pump) 2 (Auto Trap); Auto Trap / SR25 به ازای هر مبدل حرارتی. |
| ویژگیها: |  |  |  | Highlights: |  |  |  |  |  |  |  |  |
| - تنظیم دمای نقطه شبنم خروجی از +2 °C تا +7 °C - پایداری دمای تنظیم شده خروجی گاز ±0.1 °C - دو خروجی رله آلارم اعلام وضعیت آماده بکار - PID کنترلر الکترونیکی با نمایشگر دیجیتالی دما - دمای محیط کار 0….+50 °C - زمان راه اندازی کمتر 15 دقیقه و پایداری 2 ساعت - بدنه دیواری و راک استاندارد 19 اینچ - مدل ویژه جهت دمای نقطه شبنم گاز خروجی -40°C |  |  |  |  |  | - Gas outlet temperature from gas cooler adjusting from +2 °C to +7 °C - Gas outlet temperature stability ±0.1 °C - 2 relays for Status Alarm Ready to work signal - With PID electronic controller with digital display - Working Ambient Temperature: 0….+50 °C - Start-up time < 15 min and stable 2 hours in constant process condition - Wall mount Enclosure and Standard 19” Rack mount - Special Version for -40°C gas outlet dew point temperature |  |  |  |  |  |  |
| کاربردها: |  |  |  |  |  | Product Information: |  |  |  |  |  |  |
| - نصب در سیستم های آنالایزر برخط گاز در کلیه صنایع - کلیه سیستمهای آنالایزر بر خط گاز - خشک کن هوای ابزار دقیق در نقطه مصرف |  |  |  | Sample outlet dew point; Range of adjustment: +2 °C..... +7 °C, factory setting: +5 °C. Dew point stability; at const. conditions: < ±0,1°C. Sample inlet temperature; max. +180°C. Ambient temperature; 0 to +50°C. Storage temperature; -20 to +70°C. Total cooling capacity; Min 40W at 10°C … 50°C ambient temperature. Ready for working; < 15 min (stability time at constant process condition > 120 mins). Power consumption; max. 200VA (startup current at 230V=2,5A; at 120V=4,5A). Main power connection; 240V/120V, 50Hz/60Hz (a): -15%/+10%. Electrical connections; terminals: 6-12mm, cable gland: 2xPG11. Status alarm; 1 free configurable status alarm with 2 potential free change-over contacts, alarm point: < 2°C and > 8°C contact rating 250VAC, 2A; 500 VA; 50W. System of protection; IP20. Case color; white. Method of mounting; wall mounting / 19” Rackmount. Dimensions Weight; 360 x 270 x 350 mm (WxHxD) 12kg (version: 230V/115V) 483 x 270 x 350 mm (WxHxD) 17,5kg (version: 230V/115V) 19” Rackmount. Signal Output; Modbus 485 to AIEQS HMI and customer HMI. |
| Applications: |  |  |  |  |  |  |  |  |  |  |  |  |
| - Suitable to install in gas analyzer system of Cement, Steel plants - For all gas analyzer system - Instrument air drier at the using point |  |  |  |  |  |  |  |  |  |  |  |  |
| 2026 NEW Technical Information for CCA |  |  |  |  |  |  |  |  |  |  |  |  |
| -5 | -4 |  | -3 |  |  |  | -2 |  |  | -1 | How to Order | 1 |
| Note 5: Select with or without pumpSR25 or auto trap | Note 4: Choose Gas flow per each heat exchanger |  | Note 3: Select the number of heat exchanger |  |  |  | Note 2: Choose Material of heat exchanger |  |  | Note 1: Select with or without heat exchanger | CCA 992C |  |
| 0 (None) 1 (SR25 Pump) 2 (Auto Trap) | 1=150 Nl/hr 2=250 Nl/hr |  | 1/2 (Note 6) |  |  |  | 0 (None) G (Duran Glass) S (SS316) P (PTFE / PVDF) |  |  | 0 (without heat exchanger) 1 (with heat exchanger) |  |  |
| Examples: 992C1S211: CCA 992C gas cooler with 2×150 Nl/hr SS heat exchanger with SR 25 per each ÷exchanger 992C00120: CCA 992C gas cooler without heat exchanger without drain option prepared for one 250 Nl.hr heat exchanger Note 6: max two heat exchangers of 150 Nl/hr and only one heat exchangers of 250 Nl/hr is install in this type gas cooler. |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  | Condensate Removal Options | 2 |
| Automatic condensate trap |  |  |  |  | Peristaltic pump SR25.2 |  |  |  |  |  |  |  |
| 1509K3000 (optional) Automatic condensate trap with float type AD-SS. Condensate-IN: 1/2"NPTi, Condensate-OUT: 1/2"NPTi. Material: stainl. steel. |  |  |  |  | 1501P1300 (Optional) Peristaltic pump SR25.2, for panel mounting Wall thickness max.1,5mm capacity of 0,30 l/hr, pressure range 200mbar up to 2,2bar abs. Material: Novopren, PVDF, connections for tube DN 4/6mm, power: 115/230V, 50/60Hz. |  |  |  |  |  |  |  |
| توجه : مشخصات فنی دستگاه با توجه به نیاز مشتری با هر ابعاد بدنه و توان برودتی و هر تعداد مبدل در زمان سفارش مورد قبول است. |  |  |  |  | Other Technical Specification and other Enclosure Dimensions acceptable on order. |  |  |  |  |  |  |  |
| SS |  | PV |  |  |  |  |  |  | GL |  | CCA Part No Technical Information | 3 |
| - Jet-stream heat exchanger for cooler - Material: ss316 - Stagnant space approx40-100ml - Sample gas pressure abs. Max. 6barg - Made in IRAN 9910C1250; 1 Stream; 250Nl/h. 9910C2150; 2 Stream; 150Nl/h. | - Jet-stream heat exchanger for cooler - Material: PVDF - Stagnant space approx40-100ml - Sample gas pressure abs. Max. 3bar - Made in GERMANY 1593K0170; 1 Stream; 250Nl/h. 1597K0110; 2 Stream; 150Nl/h. | - Jet-stream heat exchanger for cooler - Material: Duran glass - Stagnant space approx40-100ml - Sample gas pressure abs. Max. 2bar - Made in GERMANY 1593K0140; 1 Stream; 250Nl/h. 1597K0100; 2 Stream; 150Nl/h. |  |  |

|  | Dimensions | 4 |
| --- | --- | --- |
| FRONT VIEW |  |  |
| (1); Sample Gas Inlet / Outlet; ورودی و خروجی گاز Sample; (1). (2); Status LED (compressor works OK); LED نشانگر کارکرد کمپرسور; (2). (3); PID Controller Digital Display; PID کنترلر و نمایشگر دما; (3). (4); Condensate Outlet; خروجی آب از مبدل; (4). |
| FRONT VIEW |  |  |
| (1); Sample Gas Inlet / Outlet; ورودی و خروجی گاز Sample; (1). (2); Status LED (compressor works OK); LED نشانگر کارکرد کمپرسور; (2). (3); PID Controller Digital Display; PID کنترلر و نمایشگر دما; (3). (4); SR25.2 Peristaltic Pumps; پمپ های SR25.2 جهت مکش آب از خروجی مبدل; (4). |
| FRONT VIEW |  |  |
| (1); Sample Gas Inlet / Outlet; ورودی و خروجی گاز Sample; (1). (2); Status LED (compressor works OK); LED نشانگر کارکرد کمپرسور; (2). (3); PID Controller Digital Display; PID کنترلر و نمایشگر دما; (3). (4); Auto Condensate Trap Outlet; اتو ترپ برای خروجی آب از مبدل; (4). |  |
| SIDE VIEW |  |  |